# Important parameters in SCINTILLATORS

The Inorganics are the most used, only the last one is organic (plastic) [infatti gli inorganici sono I più efficienti, e non ci possiamo permettere di non essere efficienti in nuclear medicine].

Parameters to consider:

2. specific gravity = DENSITY[g/cm^3] important because in the linear abs. coefficient (important to stop radiation) we have the "intrinsic" u' * density. (In gasses we can change the density).

More the material is dense better will stop. (plastic has like density of water.. it's bad :() [queste 3 da sapere per orale]

BGO and LSO, most used in PET (highest density) (today LYSO). In PET we have g-ray of 511keV. first req. is to stop the photon! We need a very dense and efficient scintillator.

2. DECAY TIME: "tau" of emission of scintillation. The faster the better. The number of photon in the pulse is the integral of the curve. So for the same time we have more in one with a steep curve. Also we rely on timing, so it's better to do it with something fast. E.g. comparator with threshold is easier with steep curve. Also the integration will have a rise time = to the pulse, so we get a steeper pulse also. And it's better for pile-up. Important in PET, relies a lot on timing.

3. ENERGY CONVERTED IN PHOTONS: how many ph. the scint emits for a given energy of g-ray absorbed. Efficiency reported in terms of number of photons produced by a MeV of en.

In a scint they are important b/c the statistics that dominates this value is poisson.. so we have Nph => variance = Nph.

Calculating SNR we have SNR=NphNph= Nph Larger Nph the better. We need the largest number.

So we use LSO even if CsI is better, b/c in PET the energy of the photon is A LOT!

SIGNIFICATO: il SNR dovuto allo spread intrinseco (segnale poissoniano) è importante. Vogliamo avere un SNRi buono. Significa avere un segnale alto, che significa avere una efficienza elevata. Se però consideriamo il PET, l’energia che arriva è molto grande, quindi anche con una efficienza un po’ più bassa (a favore di una densità maggiore, o di una velocità maggiore) abbiamo un segnale molto elevato.

### Generally speaking:

Timing and number of emitted photons are the reason BGO has been replaced by LSO.

In some cases we have two values of scintillator time, and a relative %. This is because some emit with a dual time constant (slide 11).

To build a SPECT: energy of Tec99 is 140keV. (photon to stop in spect). We can use CsI and NaI. Indeed using these, we have very high efficiency! CsI is the record one! :D This is important, because with 140keV we dont get a lot of photons (to resepect PET). So in this case this value is quite important. [praticamente: nella PET l’energia è così alta, che anche se l’efficienza non è altissima, I fotoni prodotti sono tanti. Nel PET ci possiamo concentrare più sul timing per esempio.]

Timing: they are not really good! but spect is not based on coincidence!! (instead PET it is!!!!)

SPECT:

- we privilege the conv. gain

- we accept a not record gravity (En. Non così grandi come nel PET difficilissime da fermare)

- we accept ok timing (non basato sulla coincidenza!)

LYSO is radioactive: but it's not a problem for PET, since a PET event is actually produced by two coincident rays. We don’t detect just one photon.

Ofc there's a small probability that we'll have 2 simulataneous event due to LYSO, that's just bad luck.

In SPECT we can't use a radioactive scint!!! B/c spect is sensitive to SINGLE PHOTON.

** PLASTIC SCINT **

The only good thing is the speed!

Then we have:

Wavelength of max emission: spectrum of emission is not like a laser one, they have a broad specturm of emission, but it's centered to a given wl. The column gives us the one of TOP emission (il picco massimo). This is important because we have to choose a PHOTODETECTOR that should cover this. tipically PMT covers the lower range, Si photodetector covers the higher range. So depending of wl. of max we'll choose one or another p.d. [cioè: va guardato lo spettro di emission per decidere che fotodetector porre dopo lo scintillator]